Battery Disposal Protocol

What to do in the event of an emergency

Covering lithium batteries and supercapacitors

- Batteries contain chemicals (that can be harmful) & stored energy (which means they can never be ‘turned off’ and therefore will always have a voltage and can cause dangerous short circuits).

- Supercapacitors contain chemicals (that can be harmful) & stored energy when charged, and have very low internal resistances, meaning they can discharge all their stored energy in a few seconds.

- Small mistakes, while working on cells, can cause explosions, fires, burns and death. When not in use, cells should be stored in ammo boxes.

- When charging, the batteries should be always supervised and kept in a mostly covered ammo box.

Typical cells used.

- There are no chemical hazards if cells are used according to their intended purpose and in compliance with the handling instructions and if the cells are not mechanically damaged. The substances contained in the cells are safely enclosed when the cells are used for their intended purpose.

- Cases of incorrect electric use (e.g. overcharge), exposure to extreme ambient conditions (e.g. high temperatures, contact with chemicals), strong mechanical loads (e.g. shock or being dropped, deformation, covering the terminals, manipulation or opening the housing and burst discs) can cause violent reactions in the cell with strong heat development and possible flames or sparks. It is possible for the materials contained in their cells and their combustion products to be emitted into the ambient air as gases and dust. Cases of incorrect use can therefore cause among other emissions of dust containing nickel/cobalt oxide powders and hydrogen fluoride or carbon monoxide fumes.

- Therefore, always refer to the manufacturers supplied information regarding the specific chemistry of the cells being used.

In the event of a battery or supercapacitor cell or module overheating

- 60 C is the safe maximum operating temperature for most lithium polymer cells, above 70 C there is risk of damage to the cells triggering runaway o o incidents. These numbers are assuming a standard lithium-ion battery, they should be checked for other battery chemistries.

- If the temperature rises above 60 C (depending upon electrolyte type) testing should be stopped or the vehicle shut down, a 1-meter perimeter o should be maintained, except for trained battery personnel who may monitor the temperature. If temperature settles, then once it has fallen below 50 C the test or vehicle can be considered safe.

- If the temperature rises above 70 C testing should be stopped or the vehicle shut down, maintain a 5-meter perimeter, monitor the temperature o remotely if possible, and if not to wait for at least 1 hour during which nothing has happened before monitoring the temperature directly. Again, if temperature settles then once it has fallen below 50 C the test or vehicle can be considered safe, but the battery modules must be removed, and o inspected and tested by trained personnel before they can be used again.

- Note that damage can be delayed, therefore never resume testing once the temperature has settled.

| Temperature rises above 60 degrees Celsius | - Stop the test. - Maintain 1 meter exclusion zone. - All to cool, monitor temperature safely. - Safe to continue once below 50 degrees Celsius. |
| --- | --- |
| Temperature rises above 70 degrees Celsius | - Stop test/Shut down test. - Maintain 5 metre exclusion zone. - Allow to cool, monitor safely. - Consider safe once below 50 degrees Celsius, but DO NOT restart testing. - Inform project supervisors. |

In the event of an overheating or short circuit event resulting in a fire

- Inform security on 020 7589 1000 (South Kensington) asking them to inform the fire brigade if the fire is large (i.e. a module and not just a cell) or if it has spread or there is a risk it will spread to other areas

- Evacuate personnel from the vicinity.

- Nobody should attempt to fight a fire unless they are trained and using self-contained closed-circuit positive pressure respirators and full safety clothing. Therefore, it is never safe for a lab user to attempt to fight a lithium battery fire under any circumstances.

- CO2 fire extinguishers can be used to cool down hot cells or battery modules if they have not ruptured or caught fire. Doing so helps to prevent 2 the spreading of fire by suppressing secondary reactions and spread the reaction over a longer period, therefore lowering peak temperatures.

- If a single or a few cells are close to catching fire but have not ruptured or caught fire yet, they can be covered in sand, vermiculite, or similar inert high temperature resistant material to suppress material releases and lower peak temperatures. This scan be found located within a designated ammo box located in your lab space.

- If necessary, and only if safe to do so, insulated bolt cutters may be used to sever cables to isolate individual cells. Individual cells which are entirely disconnected and physically isolated may be cooled with large quantities of water or water-based foam, but CO is still the preferred method.

- Water may act as an accelerant to the reaction, but if provided in significant excess then the excess water can absorb both the products of the reaction and the heat generated and eventually will cool the reaction to the point that it can be controlled or will cease. The flow of water must be continuous until all reaction has ceased and will require considerable volumes likely to damage other equipment or infrastructure, therefore should be considered as a last resort.

- Once the fire has been extinguished, ventilation to the site must be maximised to ensure reaction products from the fire have been safely removed. Personnel without the required breathing apparatus (see above) should not be allowed to enter the vicinity of the incident site until this has been done. A safe time to allow for products to dissipate is dependent on location.

- In general, if it is safe to do so and it cannot spread, it is often best to allow the chemical reaction to proceed to completion (i.e. let it burn itself out), maintaining a safe exclusion and ventilation zone, and only manage the fire in such a manner as to minimize peak temperatures.

Hazardous waste disposal of a damaged cell

- Precautions

- This procedure is NOT to be performed alone.

- Procedure is to be done under adequate ventilation.

Additional PPE to be used during procedure:

- Nitrile gloves

- Breathing mask fitted with organic vapour filter (if dealing with cells that have failed)

- Purpose of Procedure

- This is a general procedure for disposing any lithium-Ion cells.

- Required Equipment

- Sand or vermiculite.

- An Ammo box.

- Nitrile gloves

- Disposal procedure

- Raise a hazardous waste number. (This step can be done after filling all the boxes, if there is a significant quantity of waste, and you think it may take some time to complete the task)

- Go to http://www.imperial.ac.uk/estates-facilities/buildings/services/waste-disposal/waste-directory/hazardous/ and navigate to the online defect form.

- click "Raise a soft service request" and fill out the relevant sections:

- Property:  *enter building name*

- Request type: Hazardous waste collection

- Is halls of residence: No

- Access restrictions: No

- Click submit; an email will be sent to your imperial address from Customers service centre with the waste number as the subject (can take a few hours)

For cells that have previously been deemed dangerous (e.g. placed in ammo boxes)

- ensure you are wearing all required PPE: Lab coat, glasses, nitrile gloves and respirators with filters.

- Open the ammo box (cells should be stored in the box in bags or the terminals covered, if not do so as you remove the cells from the box)

- Fill thin layer of vermiculite or sand in ammo box and place cells onto vermiculite or sand.

- Record the type and number of cells that have been transferred into the ammo box.

- Cover the layer of cells with more vermiculite or sand.

- Repeat 4 & 5 until box is full.

- Email Roger Smith (roger.smith@imperial.ac.uk) with subject "Hazardous waste collection: No. (Waste number)" and attach the hazardous waste form.

- Roger will allocate you a time slot, at this time take the blue boxes to the chemical waste shed. The shed is located in a large blue prefab at the back of the Blackett Building.

General Cell Storage Procedures

Cells must be stored as follows:

- In a locked and secure environment, this must be labelled as containing charged cells which are dangerous if handled incorrectly and only authorised users are permitted access, ‘What to do in the event of an emergency’ must be clearly displayed nearby

- Access must be limited to approved users, and a log maintained of who has access

- Cell must be stored ideally with a controlled temperature and humidity below 10 C but at least below 20-25 C and in a non-humid environment

- Cells must be stored such that there is no chance of a short-circuit between the two terminals of a cell or between multiple cells

- Cells must be stored such that there is no chance of mechanical damage to the cells during storage or handling

- Cells must never be left in the open unsupervised

All cells must be considered to be ‘charged’ unless otherwise established:

- Batteries are always ‘charged’ unless verified and clearly identified as dead

- Supercapacitors must be considered ‘charged’ unless a yellow or orange short circuit wire is securely attached to both terminals

- Attaching the short circuit wire directly to a charged supercapacitor (even with a relatively low voltage) will cause high currents to flow instantaneously through the wire potentially vaporizing it and causing burns or fire

Toxic and dangerous chemicals can be released in the event of thermal runaway or explosion, both solids, liquids and gases, and large amounts of heat energy can be released very quickly, therefore:

- Cells should never be stored where pressure can build up to dangerous levels, Work must be carried out in well-ventilated areas Access and escape routes must be thought out in advance and access to escape routes maintained at all times Any other users in the same ventilated area as the ‘working area’ must be aware the work that is going on and the risks associated with it, the evacuation route, and ‘what to do in the event of an emergency